## Slide 1: Development of the mid-IR Echelle high-dispersion spectrograph employing the germanium immersion grating
Yasuhiro Hirahara*a, Tsuyoshi Hiraoa, Yoshio Tatamitania, Tomohisa Yonezua, Noboru Ebizukab, Kentarou Kawaguchic, Hitoshi Tokoroa, Tomomichi N. Okaa
aDepartment of Earth and Planetary Sciences, Graduate School of Environmental Studies,
Nagoya University, Nagoya 464-8602, Japan
bDepartment of Electrical Engineering and Computer Science, Graduate School of Engineering, Nagoya University, Nagoya 464-8603, Japan
cDepartment of Chemistry, Okayama University, Okayama 700-8530, Japan
Notes: We have developed a germanium immersion grating mid-infrared cryogenic spectrograph (GIGMICS) designed for the Nasmyth focus stage of NAOJ Subaru 8.2 m telescope, which operates at N-band (8-13 mm) in wavelength with the R ~ 50,000. A single crystal germanium immersion echelle grating (30 ﾗ 30 ﾗ 72 mm) for collimated beam size of f28 mm was fabricated by utilizing ultra precision micro-grinding method coupled with the ELID (ELectrolytic In-process Dressing) technique (Ohmori, H. 1992)1. All optical components are arranged on the 800 mm diameter cold optical base plate (~30 K) of the cryostat. By the Si:As IBC (Impurity Band Conductor) focal plane array (FPA) detector (412 ﾗ 512 pixels, unit pixel size 30 mm) operated at 5 K simultaneously acquires ~13 % wavelength coverage for N-band. The instrument has been assembled and is now tested for the application to the gas-phase IR high-resolution spectroscopy.

## Slide 2: Interstellar Molecules
  ○Simple Hydrides, inorganic species：
  H2, HF, HCl, H2O, NH3, C2, O2, CH4, CO, CO2, N2O, H2S, CS, SO2, OCS, SiH4, SiO, SiS, NaCl, KCl, AlCl, AlF, PN, HCP
  ○molecules in cyclic (AROMATIC) form：
  c-C3H, c-C3H2, c-(CH2)2O, c-C3H2O(cycloprpenone), C6H6 (benzene), c- SiC2, c- SiC3,
  ○”Cyanopollynes, polyacetylenes and related “LINEAR Carbon chains”：
  C3, C5, C3O, C3S, C3O, C3S, H2C3, H2C4, H2C6, C2H2, HC4H(diacetylene）, HC6H(triacetylene), C2H4, C4Si, HCN, HNC, HC3N, HNCCC, HCCNC, HC5N, HC7N, HC9N, HC11N, CH3CN, CH3NC, CH3CH2CN, CH3C3N, CH3C5N, CH3CCH, CH3C4H, CH3C6H, C2H5CN, C2H3CN, C2H5CN
  ○Ions：
  H3+, CH+(OPT), CO+, SO+, CF+, HCO+, HOC+, HN2+, HCS+, H3O+, H2COH+, HOCO+, HCNH+, HC3NH+, C4H−, C6H−, C8H−, C3N-, C5N-
  ○Radicals：
  CH, CH2, CH3, OH, NH, NH2, SH, HNO, SO, NS, NO, SiC, SiN, NaCN, MgCN, MgNC, AlNC, SiCN, SiNC, NH2CH, HCO, CCH, C3H, C4H, C5H, C6H, C7H, C8H, CN, C3N, C5N, H2CN, HCCN, HC4N, CH2CN, CCO, CCS, CP, PO
  ○Aldehydes, Alchoholes, Ethers, Ketons, Amides and related species(“Pre-Biotic” molecules)：H2CO, H2CS, CH3CHO, HNCO, HNCS, NH2CHO, HC2CHO, CH2OHCHO, CH3OH, C2H5OH, CH2CHOH, CH3SH, (CH3)2O, (CH3)2CO, HCOOH, HCOOCH3, CH3COOH, H2CCO, CH2CCHCN, CH2NH, CH3NH2, NH2CN, CH3CONH2
Detected in IR
Only(≒Nonpolar)
    Few IR “line survey” in IR, “No” laboratory IR data
    Orion KL, 4-25 mm, IRTF (3.0 m) & IRSHELL (Evans et al. 1991)
    IRC+10216, 8-13 mmKitt Peak Mayall telescope (3.8 m) & FTS (Keady & Ridgway 1993, Boyle et al. 1994)
C60, C70
(Spitzer)
, H2O+, H3O+, OH+, SH+, H2Cl+(Herschel)

## Slide 3: Science in N-Band
“All molecules have Vibrational transition”
CH4, C2H6 , …, Carbon Chain, “PAH”s
(SiO)x ….
Chemistry Related to Star Formation
“Fingerprint” region for Complex Organic molecules: ν=2π√(k/μ), k: force constant
N-Band 7.5 ~ 13.5 mm
Atmospheric Opacity
Notes: How important 10 micron region is!
Fingerprint region.
Higher resolution, more informative.

## Slide 4: Ion molecule reaction network for Carbon
Key ions:
H3+, CH3+
C3H3+: “Smallest Aromatic ion”: n(π e-)=2

## Slide 5
“Rotationally resolved” vibration spectrum is useful for
○definite identification of a “new” molecule
○accurate derivation of abundances and excitation
○the understanding the “chemistry”=formation pathways
R=10,000-200,000
Required Spectral Resolution

## Slide 6: Germanium Immersion Grating
Resolving Power, R = l/Dl = DL/l = 2L/l
DL=2L
Conventional Grating
L
Grooves on the Ge Surface : Fabricated by RIKEN’s ELID (ELectrolytic In process Dressing) Micro-machining Method (Ebizuka et al., 2003)
m=0.6 mm
β= 68.75º
72 mm
30 mm
30 mm
N = 130 grooves
L
n
DL=2nL
Immersion Grating
L
n
→ Compact Cryogenic Optics
Resolving Power, R = l/Dl = DL/l = 2nL/l (n ~ 4 for Ge)
Surface Roughness:
~3 nm
R~51,000 @ 10μm
For 25mm beam diameter
Notes: The key device for our new spectrometer is the Ge immersion grating, which is directly grooved onto the surface of wedge shape of single crystal Ge which is transparent but highly refractive material. In comparison with the conventional grating, the effective path-difference delta f increases by factor of the reflactive index n, so that the size of the grating becomes compact. making fine and precisely periodical grooves on the hard and fragile Ge surface is overcome by the great effort of Dr, Ebizuka and Prof. Ohmori’s group, by adopting their unique machining technique ELID - Electrolytic In-process dressing micro machining method. As Size dimension is shown here, this precious devices is used for our cryogenically cooled echelle grating with very high diffraction order.

## Slide 7: Optical design of GIGMICS
Slit width 0.612 arcsec (0.488 mm) at R = 50,000, F/12.5
Echelle grating Ge immersion grating
Groove interval : 0.6 mm
Blaze angle : 68.75 deg.
Order : 344-560
Cross Disperser Short wavelength (7.5-10 m)
Groove constant : 100 grooves/mm Blaze angle : 26.7 deg. Long wavelength (10-13.5 m)
Groove constant : 61.97 grooves/mm Blaze angle : 18.1 deg.
･ Collimator Al spherical mirror Focal length : 500 mm (F/20)
･ Camera Ge and ZnSe lenses Focal length : 192.5 mm (F/7.7)
･ Detectors 512 x 412 Si:As IBC (Spectrograph) 5K cooled
256 x 256 InSb (Slit viewer) 30K cooled
(Cryostat x 2, for 5K, 30K for optics, 70K for Radiation shield)
･Cryogenic actuators 4 (Filter Slit, Cross disperser, Focus)
Notes: I will breafly show you the optical layout, which is a basic lithorow type . 2 detector are mounted on the disk of 80cm diameter, one for imagingm and the other for taking echelle format. The slit, immersion grating. Repetively turn down

## Slide 8: Optical layout of GIGMICS
Immersion Grating
Camera Lens (F/7.7)
Si:As
512*412 pixel, 6 K
Cryostat: f800, 30 K
Collimators (F/20, f80)
Cross Disperser
Slit Mirror (0.488 mm)
InSb (Slit Viewer)
256*256 pixel
Fore-optics
SP-optics
＊Before mounting filter and slit turrets
K-band image of Saturn
Mitaka 1.5mφ, F/12.5
(Subaru simulator)
Notes: This picture shows the optical setup of GIGMICS, consisting of fore-optics and spectrograph optics, described in detail.
  1.1 Fore-optics
Light entering the cryostat at f nunmber/12.5 through the window is collimated by the first spherical mirror (SM1). Through the cold Lyot stop and the BaF2 blocker (F1).
The second spherical mirror (SM2) reimages the telescope’s focal plane at f/20 onto the slit mirror,
to reflect a 8arcsec beam on an 8m telescope field of view into the slit viewer shown here,
used to guide on an astronomical object at the slit. here is illustrated the result of K-band image of saturn.
In this region, filter and lens turret is mounted, although not shown in this picture. Moreover, this slit mirror module has now been replaced with slit turret. Note that the slit size is appropriate for the point source diffraction limit for the telescope for10 micron.
==MOTION==
  1.2 Spectrograph optics
The Light passing through the slit enters into the Littrow-mount echelle spectrograph.
The light reflects off the flat fold mirror through the long pass filter, which cuts on at sevenpoint five 7.5 mm,
and then expands at f/20 to the spherical mirror here (SM3).
The mirror collimates the light, creating a 25 mm diameter beam on the Germanium immersion grating (IG).
The dispersed light returns to SM3,
and then go to the second spherical collimating mirror (SM4),
which reflects it to the cross-disperser (CD) mounted on the
cryogenic rotational stage driven by the worm gear mechanism.
The cross dispersed light is reflected to the spectrograph detector here,
through the Ge - ZnSe achromatic doublet lens (L3+L4).
The lenses is mounted on the flex stage shown here for focus adjustment.
The focal reducer changes the beam to f/7.7,
to satisfy the Nyquist sampling condition on the detector for the diffraction pattern.
==END==
It is used in very high order, ranging from 332 at 13.5 mm to 597 at 7.5 mm.
The groove facets are tilted 0.69ｰ relative to the incident aperture of the immersion grating to eliminate the ghosts caused by reflection from the incident aperture.
The camera lens images the focal plane at 1.15 : 1 demagnification onto D1.
The ZnSe collimating lens (L2) between the Lyot stop and the second collimating mirror allows the slit viewer to be used for Lyot stop viewer.
The Lyot stop viewer is useful for alignment with the telescope optics. A set of the collimating lens (L2)
and the second collimating mirror (SM2) forms images of the Lyot stop and the pupil onto the slit with a scale increase of 1 to 1.06.
The camera lens of the slit viewer transfers these images,
( It is composed of the ZnSe camera lens, the long pass filter, which cuts on at 2.75 mm, and the detector, 256 ﾗ 256, 30 mm/pixel InSb FPA detector (D1). The collimating lens is mounted on the lens turret (LT), so as to remove from the optical pass during spectroscopic observations.)

## Slide 9: Motions(1): Focus adjuster & Crossdisperser stage
Bending of YH75 (t=0.5 mm)
Focal lens mount stage:+/- 1.0 mm, Resol. 0.01 mm
Flexture stage with parallel spring
Spur gear
Commercial stepper motors (Escap P 430-258-005), bearings of which are replaced with MoS2 sputtering bearings of NTN Corp.
(Z-Lab, Nagoya Univ.)
Cross Disperser

## Slide 10: Motions(2): Filter Turret & Slit Turret

## Slide 11: Echelle Spectral Formats

## Slide 12: Cryostat
2 Cryocoolers
Outer radiation shield: ~70K
Inner radiation shield,
Optics on 800mmφStage,
& Slit viewer(InSb FPA detector)
: 3０K
Si:As BIB FPA detectpr: 5~7K
Notes: GIGMICS has a 30 K cooled inner optical chamber surrounded by an outer chamber that is cooled down to 80 K, all contained in a large vacuum jacket (Fig. 4). The vacuum jacket weighting about 1,000 kg consists of a cylinder-shaped base section onto which a large O-ring sealed covers are bolted. The base section and the cover are made of stainless steel (SUS 304). The external dimensions of the vacuum jacket are 1010 mm diameter ﾗ 1250 mm height. The G10 fiberglass supports are used to separate the outer chamber from the vacuum jacket wall and the inner optical chamber from the outer chamber. The outer chamber is wrapped with twentyfold layered aluminized Mylar sheets to reduce the radiation from the inner wall of the vacuum jacket.
GIGMICS is equipped with two closed cooling systems. A two-stage cryocooler (RDK-408D: Sumitomo Heavy Industry Ltd.) cools the optical chamber and a Si:As IBC FPA detector down to 30 K and 4 K, respectively. The first stage has a rated cooling of 37 W at 40 K, and the second stage has a rated cooling of 1.0 W at 4.2 K. A single-stage cryocooler (TZ-65: Daikin Co.) with a capacity of 60 W at 80 K cools the outer chamber down to 80 K. At the beginning of cooling, the cooling speed is boosted by pouring LN2 into pipes attached to base plates of both chambers.
In order to minimize the deviation of optical axis caused by the difference of thermal contraction, the optical base plate and most of optical components were fabricated from the A6061 aluminum alloy.

## Slide 13: Echellegram of NH3 in Emission at 370 K
n2 (inversion)
820 cm-1 (12.2 mm)
1,100 cm-1 (9.1 mm)
By Rotating Cross Disperser(s), Observable Spectral Region between 7.5 mm and 13.5 mm Can Be Selected.
Liq. N2 Cryostat (110 K) Background Reducer

## Slide 14: Comparing Echellegram with FT Spectra
º
FWHM = 0.03~0.05 cm-1
R~30,000

## Slide 15: Laboratory GIGMICS spectrum (II): methane
q=32º （Cross Disperser Angle) 31º
1390 cm-1(7.2μm)
1220 cm-1(8.2μm)

## Slide 16: Laboratory GIGMICS spectrum (III): allene
(3)アレン
700 cm-1(14.3μm)
950 cm-1(10.5μm)
= 188 193
H2C=C=CH2

## Slide 17: GIGMICS now at KANATA 1.5-m telescope, Higashi-Hiroshima Observatory, Since Dec. 2010 for the first light
Targets:
Vib.-rot. Transitions of Methane, Ethane, Ammonia, N2O, O3, SO2, H2O, CO2, SO2, H2S, NOx, Halogen Oxides, etc.,
in the Planets, Stellar Atmosphere, bright SFRs, CSE of late type stars, and the upper atmosphere of the Earth
End：

## Slide 18: Supl: Sensitivity
Nphoton(Dark) = 150 photons·pixel-1·s-1 : Good!
Observed Ratio of “Signal-to-Noise Ratio (SNR)”
(SNR)Ech/(SNR)FT = 100
Ratio of Theoretical SNR
  for Single-Channel System (Detector Noise Dominant),
(SNR)G/(SNR)FT = [N]-1/2 ~ 0.002
  for Multi-Channel Grating Spectrometer (Echellegram)
(SNR)Ech /(SNR)FT = [N]-1/2 · Neff = 100 (Neff ~ 50,000)
  If Modulation Noise Dominant
[(SNR)Ech /(SNR)FT] = [N]+1/2 · Neff ~ 2 ×107
    N = 512 ×412 = 210,944 (Number of Spectral Elements).

## Slide 19: Echellegram of CO2 Laser with Thermal Radiation
0.000 sec
Nphoton
0
l
m
0.667 sec
10.59 mm m=414
10.57 mm m=415
10.55 mm m=416
10.59 mm m=415
425
420
415
410
405
m = 400
Dispersion : dl ~ 8 ×10-5 mm·pixel-1 (~0.007 cm-1·pixel-1) @ 10.6 mm
0.67 mm (58 cm-1)
~10 % of Bandwidth in One Integration
ml = nd[sin(b) + sin(g)]
11.03 mm (907 cm-1)
10.36 mm (965 cm-1)

## Slide 20: First-Light Observation with Slit Viewer
Diameter：1.5 m
F：12.2
Field of View：15 arcsec (88.7 mm/arcsec)
Diameter：1.5 m
F：12.2
Field of View：15 arcsec (88.7 mm/arcsec)
Mitaka 1.5m Telescope, March – May 2005
Slit Viewer Performance : ~50% of Estimated
- Limiting Magnitude : 7.1 mag. at K-band (2.2 mm)
(1.0 sec exp. time, S/N = 10, seeing = 1”.0, Nreadout = 620e-)
Mitaka 1.5mφ, F/12.5
(Subaru simulator)

## Slide 21: Signal Detectability of IRHS
Nphoton(Dark) = 150 photons·pixel-1·s-1
Nphoton(Background Radiation) :
  Troom = 290 K : 1,500 photons·pixel-1·s-1
  Teff = 210 K : 250 photons·pixel-1·s-1
cf. Full-well Capacity (Si:As, 5 K)
~ 200,000 photons·pixel-1 (t = 500 sec for Teff = 210 K)
Detectable If Nphoton(Signal) > [Nphoton(BR)]1/2
= 16 photons·pixel-1·s-1 (Teff = 210 K)
Cryogenic Cooling Allows Us to Extend Integration Time & Enhance Signal-to-Noise Ratio.